## Slide 1: Aerogel mass production for the CLAS12 RICH: Novel characterization methods and Optical Performance
RICH 2016, September 9th 2016, Bled
Contalbrigo Marco
INFN Ferrara

## Slide 2
The CLAS12 RICH
2
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled
CLAS12
RICH goal: p/K/p identification from 3 up to 8 GeV/c and 25 degrees
~4s pion-kaon separation for a pion rejection factor ~ 1:500

## Slide 3
3
Contalbrigo M.
The Hybrid Optics Design
plane
mirror
spherical
mirror
photon detector
aerogel
2 cm
6 cm
gap
p
g
plane
mirror
spherical
mirror
photon detector
aerogel
2 cm
6 cm
gap
p
g
Minimize active area (cost) to about 1 m2
Material budget concentrated where TOF is less effective
Focalizing mirrors allow thick radiator for good light yield
Time resolution < 1 ns to distinguish direct and reflected patterns
Direct rings and best performance for high momentum particles
Reflected rings for less demanding low momentum particles
4.2 m
3.7 m
1.2 m
p
Beam
RiCH 2016, 9th September 2016, Bled

## Slide 4
Aerogel Production
4
3 + 3 cm
2 cm
44 x 2 tiles
32 tiles
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled
72 (3 cm thick ) and 22 (2 cm thick) full squared tile + 30 shaped ones
High transparency and large refractive index (n=1.05) to ensure photon yield
Large area 20 x 20 cm2 to reduce losses at the edges, variable thickness (2 and 3 cm)

## Slide 5
Aerogel Radiator
5
Flexible geometry, mass production capability
Achieved ~ 0.0050 mm4 cm-1 clarity for large tiles (LHCb had 0.0064 mm4 cm-1 for n=1.03)
Collaboration with Budker and Boreskov Institutes of Novosibirsk
Dominated by
Scattering length
Hygroscopic aerogel requires special care and dry N2 atmosphere
A0 > 95 %
<Lsc> ~ 50 mm
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled
Dominated by
Absorption coeff.

## Slide 6
Aerogel Characteristics in the Air
6
Contalbrigo M.
Transmission vs Time
Absorption and Clarity @ 400 nm
Monitoring the time dependence of the transmission of aerogel tile
in environment of non-zero relative humidity (~ 40 %)
RiCH 2016, 9th September 2016, Bled

## Slide 7
Aerogel Storage
7
Contalbrigo M.
Within sealed envelops and inside a dry-cabinet (few % RH)
RiCH 2016, 9th September 2016, Bled

## Slide 8
Aerogel Weight
8
Contalbrigo M.
In Air ( ≈ 40 % RH)
Inside Dry-Box ( ≈ 1 % RH)
RiCH 2016, 9th September 2016, Bled
(hours)

## Slide 9
Aerogel Specifications
9
Contalbrigo M.
OPTICAL:
Density 0.223 < r < 0.245 gr/cm3
Refractive index (n2=1+0.438 r) 1.0477 < n < 1.0523
Scattering length Lsc > 43 mm
Absorption coefficient A>0.95
MECHANICAL:
No bubbles, crackes; chips limited to less than 1 % area
Side to side length variation DLside < 0.25 mm
Tile to tile thickness variation DHtile < 1.5 mm
Surface planarity DSsurf < 1 % of lateral side
RiCH 2016, 9th September 2016, Bled

## Slide 10
Mechanical Performance
10
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled
From diamond wire to diamond wheel cut to
improve speed and precision, reduce defects (cracks)
Tendency to bend during baking procedure
(significant fraction of rejected tiles)

## Slide 11
Optical Performance
11
<Lsc> ~ 50 mm
A0 ~ 97 %
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled
In line with R&D achievements
Notes: Cercare foto di Tayler o nostra con aerogel

## Slide 12
Optical Performance
12
<Lsc> ~ 50 mm
A0 ~ 97 %
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled
BELLE-II
See talk of R. Pestotnik on Monday
In line with R&D achievements
Notes: Cercare foto di Tayler o nostra con aerogel

## Slide 13
Refractive Index
13
Spread in refractive index corresponds to ~ 3 mrad of Cherenkov dispersion
Tile by tile index of refraction will be accounted for into the reconstruction
program to suppress the spread in refractive indexes (Cherenkov angles)
RMS ~ 3 mrad
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled

## Slide 14
Resolution | Direct (mrad) | Reflected (mrad)
Emission Point | 1.7 | 1.7
Readout Accuracy | 2.1 | 1.0
Chromatic Aberration | 3.0 | 2.5
Aerogel Optical Prop. | ≤ 1 | ≤ 2
Mirror System |  | ≤ 1
sq (1 p.e.) | 4.2 | 3.9
Requirements | Direct | Reflected
Max. momentum | 8 GeV/c | 6 GeV/c
sq (4s separation) | 1.4 mrad | 2.5 mrad
Np.e. Yield | ≥ 10 | ≥ 3
CLAS12 RHIC: Resolution
14
Contalbrigo M.
Np.e.
Validated with large scale prototype
Studied in laboratory
RiCH 2016, 9th September 2016, Bled

## Slide 15
Aerogel Chromatic Dispersion
15
Contalbrigo M.
Measured by prisma method:
Measured by prototype with optical filters:
Expected value from density:
n(400nm) = [1+0.438r]1/2 = 1.0492
RiCH 2016, 9th September 2016, Bled
Notes: Qual e’ il packing factor di HERMES ?
Refraction or refractive ?

## Slide 16
Aerogel Surface Quality
16
Contalbrigo M.
Refraction from a surface
with local normal deviation q
Contribution on light dispersion
at small incident angles
sqaer
RiCH 2016, 9th September 2016, Bled

## Slide 17
Aerogel
Laser
Screen
Scan of aerogel surface
  . . . . . . . . . . . . .
  . . . . . . . . . . . . .
  . . . . . . . . . . . . .
  . . . . . . . . . . . . .
  . . . . . . . . . . . . .
  . . . . . . . . . . . . .
  . . . . . . . . . . . . .
  . . . . . . . . . . . . .
Distributions of X & Y positions of the spot
Aerogel Surface Scan
17
Contalbrigo M.
x-y axis movable table
CCD camera [ThorLabs DCU 224c]
- sensitive area [5.95-4.76 mm]
- resolution [1280-1024 pixels]
- pixel size 4.65 mm
screen
motor
laser
CCD camera
RiCH 2016, 9th September 2016, Bled

## Slide 18
Aerogel Surface Planarity
18
Contalbrigo M.
From laser spot shits to surface gradients
From surface gradients to surface map by linear regression
[rad]
[rad]
Consistent with Russian vendor planarity evaluation
Validated with touch machine measurements
RiCH 2016, 9th September 2016, Bled

## Slide 19
Surface Measurement Validation
19
Contalbrigo M.
RiCH 2016, 9th September 2016, Bled
Possibility to derive the thickness profile

## Slide 20
Light Dispersion vs Aerogel Surface Quality
20
Contalbrigo M.
Aerogel
Laser
Screen
Aerogel
Laser
Screen
Mirror
s ~ 0.9 mrad
s ~ 16 mrad
Acceptable light dispersion even
for poor planarity DSsurf = 1.25 %
RiCH 2016, 9th September 2016, Bled

## Slide 21
[mm]
[mm]
Aerogel Surface Map
21
Contalbrigo M.
Light dispersion as a function of incident angle:
RiCH 2016, 9th September 2016, Bled

## Slide 22
30 mm
CCD Camera
Laser Beam
R
Description of the setup
Take the average X & Y profiles of the spot
Forward Scattering
22
Contalbrigo M.
Analysis steps:
- Reference beam profile taken without aerogel
- Extract laser beam profile and compare
with reference measurement
- Extract angular dependence of light intensity
after passage through the aerogel
Scattering of the light in the medium due to
the anisotropy of the dielectric properties
caused by density microscopic fluctuations
RiCH 2016, 9th September 2016, Bled
s = 0.170
s = 0.180
s = 0.716

## Slide 23
Angular dependence of the measured intensity:
Forward Scattering
23
Contalbrigo M.
Differential of the measured intensity:
Negligible scattering
at angles relevant for Cherenkov resolution
RiCH 2016, 9th September 2016, Bled
Typical fraction of light above 2 mrad ≤ 1 %
Stack

## Slide 24
✔ Storing (in dry cabinets) and handling procedures
has been defined for the hygroscopic material
✔ Steadily improvement of the process
✔ Non invasive techniques employed to
verify light propagation into the material
✔ Tools for measurements and monitoring
the aerogel characteristics show stable
performance in line with the project design
Conclusions
24
Contalbrigo M.
Aerogel mass production for CLAS12 RICH is ongoing
RiCH 2016, 9th September 2016, Bled